## Slide 1: Experience with Cavity Performance: Changes from Vertical to Cryomodule Test
Dan Gonnella
Nicholas Valles, Matthias Liepe
Cornell University

## Slide 2: Motivation: The Cornell ERL
CESR
LINAC-A
LINAC-B
Injector
Beam Stop

## Slide 3: Purpose of HTC Tests
Main Linac Cryomodule
Horizontal Test Cryomodule (HTC)

## Slide 4: Cavity Fabrication and Vertical Test
Cavity preparation:
Bulk BCP + 650oC Heat Treatment + Light BCP + 120oC Bake + HF Rinse
Fabricated with 85% field flatness
CMM results show that we achieved ¼ mm shape precision after welding
Tuned to 95% field flatness
6 additional cavities have been fabricated.
String Assembly for the prototype main linac cryomodule has begun.

## Slide 5: Vertical Test Results
Chart: 
| 11.42 | 4.8099999999999996 | 5.56 | 6.36 | 7.2 | 8.1999999999999993 | 9.26 | 10.25 | 11.58 | 12.96 | 13.61 | 14.33 | 14.84 | 15.64 | 16.690000000000001 | 17.63 | 19.48 | 20.52 | 21.47 | 23.7 | 25.36 | 24.97
ERL7-1 | 16715000000 | 15613000000 | 16064000000 | 16587000000 | 16965000000 | 17039000000 | 17432000000 | 17374000000 | 17092000000 | 16841000000 | 16654000000 | 16506000000 | 16321000000 | 16161000000 | 15900000000 | 15827000000 | 15431000000 | 15294000000 | 14950000000 | 14153000000 | 13339000000 | 13512000000
ERL7-2 | 18106000000 | 18483000000 | 18263000000 | 20068000000 | 17860000000 | 19885000000 | 19314000000 | 19427000000 | 19412000000 | 17668000000 | 18228000000 | 18452000000 | 18553000000 | 17597000000 | 17757000000 | 16745000000 | 16875000000 | 16799000000 | 16009000000 | 15637000000 | 15297000000 | 15091000000 | 14972000000 | 14688000000
ERL7-3 | 26269000000 | 26220000000 | 27620000000 | 27877000000 | 29502000000 | 28730000000 | 28800000000 | 28938000000 | 28352000000 | 28223000000 | 28116000000 | 27968000000 | 28408000000 | 28882000000 | 27743000000 | 28816000000 | 28321000000 | 27594000000 | 25949000000
ERL7-4 | 25773000000 | 26469000000 | 26421000000 | 26795000000 | 27344000000 | 28209000000 | 27615000000 | 27487000000 | 27505000000 | 27262000000 | 27142000000 | 26281000000 | 25349000000 | 25068000000 | 24451000000 | 24256000000
ERL7-5 | 30423000000 | 29350000000 | 30045000000 | 31300000000 | 30737000000 | 31500000000 | 32000000000 | 31000000000 | 30800000000 | 30000000000 | 27800000000 | 27700000000 | 27100000000 | 27000000000 | 27000000000 | 26000000000 | 26500000000 | 25800000000
ERL7-7 | 28800000000 | 29904000000 | 27682000000 | 29200000000 | 29184000000 | 30200000000 | 31038000000 | 31683000000 | 32090000000 | 31448000000 | 30572000000 | 30490000000 | 29959000000 | 29551000000 | 29283000000 | 30000000000 | 30421000000 | 29549000000 | 29853000000
Q0
Eacc [MV/m]
1.8 K
Plot courtesy of F. Furuta
ERL7-1 achieved Q0 of 1.7x1010 at 1.8 K and 16 MV/m in vertical test.
Note: HF rinse was added to the recipe after the first vertical test.
ERL Spec
ERL7-2 has been reprocessed

## Slide 6: 3 HTC Tests
HTC-1: Follow vertical assembly procedure as closely as possible
HTC-2: Include side mounted, high power RF input coupler
HTC-3: Full cryomodule assembly-high power RF input coupler and beam line HOM loads

## Slide 7: HTC Assembly

## Slide 8: HTC-1 Results
Comparison of Vertical and HTC-1 test of ERL7-1
HTC-1 Results
Cavity received only HPR and helium tank welding between vertical and horizontal tests.
Q0 significantly increased in the cryomodule compared with vertical test.
Thermal cycling caused Q0 to increase.
Note: all cool downs are slow by vertical test standards: ~1 K/hour
Cool Down

## Slide 9: HTC-2 Results
Cavity again met specification in the cryomodule.
Large amounts of radiation (10 kR/hr) were seen at higher fields.
Input coupler installation resulted in a field emitter that limited cavity Q0.

## Slide 10: HTC-3 Results
Initial Cooldown at 16.2 MV/m
  Q0(2.0 K) = 2.5 x 1010
  Q0(1.8 K) = 3.5 x 1010
  Q0(1.6 K) = 5.0 x 1010
10 K thermal cycle at 16.2 MV/m
  Q0(2.0 K) = 3.5 x 1010
  Q0(1.8 K) = 6.0 x 1010
  Q0(1.6 K) = 10.0 x 1010

## Slide 11: Conclusions
Cavity performance in the cryomodule exceeded performance in vertical test.
Measurements in horizontal test cryomodule demonstrated world record quality factors:
  Q0(1.6 K, 16.2 MV/m) ≥ 1011
  Exceeded 1.8 K spec by 3X – aimed for 2x1010, got 6x1010.
10 K thermal cycling in cryomodule resulted in a 3X reduction in residual resistance (3.20->1.34 nΩ).
We’ve demonstrated that the Cornell ERL Q0 specification of 2x1010 at 1.8 K and 16 MV/m can be repeatedly met in a cryomodule.