## Slide 1: Vertical and Horizontal High Q Testing – Lessons Learned for LCLS-II
Dan Gonnella
Cornell University
October 29, 2015

## Slide 2: LCLS-II High Q R&D
3 Labs participated in R&D on High Q: Cornell, Fermilab, and JLab
This R&D consisted of:
  Design of N-Doping recipe for Single and 9-Cell cavities
  Testing of cavities vertically
  Testing of fully dressed cavities horizontally
This work was meant to demonstrate feasibility and repeatability of N-doping for LCLS-II

## Slide 3: Cornell Results: Single-Cells
5 single-cells given same doping: 800oC, ~60 mTorr N2, 20 minutes + 30 minute anneal, followed by different EP (6-30 µm)

## Slide 4: Cornell Results: Single-Cells
Average Q0 at 2K: 3.6 x 1010 at 16 MV/m
Average max field:
27 MV/m
Far exceeding LCLS-II specs

## Slide 5: Cornell Results: 9 Cells
Average of Cornell N-doped 9-cell cavities
Q0 at 16MV/m = (3.0 ± 0.3) x 1010
Quench field = 18 ± 3 MV/m
No field emission
20N30 doping: 20 min N-doping + 30 min anneal at 800C
6N6 doping: 6 min N-doping + 6 min anneal at 800C

## Slide 6: Effect of N-Doping on Quench Field
Chart: Eacc [MV/m]
| 13.04 | 11.67 | 10.56 | 9.89 | 8.73 | 7.58 | 7.1499999999999986 | 6.43 | 5.1199999999999974 | 4.24 | 3.24 | 14.84
4th final | 27400000000 | 30430000000 | 31487000000 | 32646000000 | 32603000000 | 31496000000 | 31426000000 | 31156000000 | 29843000000 | 29380000000 | 28448000000 | 19480000000
3rd VT | 26800000000 | 28400000000 | 27900000000 | 26200000000 | 25500000000 | 25170000000 | 26800000000 | 25400000000 | 22700000000 | 24490000000 | 23300000000 | 23570000000 | 22480000000 | 20978000000 | 20960000000 | 21020000000 | 20140000000 | 19840000000 | 19280000000 | 17980000000 | 17000000000 | 16300000000
2K
Baseline, undoped
2N2-doped
-7 MV/m
Same quench location before and after doping
30% drop in quench field after doping
Doping model predicts 30% drop in lower critical field after doping (~130 mT => ~90 mT)!
Should expect: 32 MV/m XFEL quench field average => 22 MV/m LCLS-II doped cavity average

## Slide 7: Effect of N-Doping on Quench Field
Chart: Eacc (MV/m)
| 1.5 | 2.2000000000000002 | 3.03 | 3.48 | 4.03 | 4.6100000000000003 | 4.93 | 5.89 | 6.84 | 7.68 | 8.39 | 8.92 | 10.02 | 11.14 | 11.84 | 13.18 | 14.36 | 15.18 | 15.67
heavy dope | 23178000000 | 23537000000 | 24859000000 | 25430000000 | 26186000000 | 27496000000 | 27652000000 | 28777000000 | 28401000000 | 28668000000 | 29056000000 | 29590000000 | 29976000000 | 31090000000 | 30187000000 | 30872000000 | 30285000000 | 30222000000 | 30118000000
baseline, none dope | 15827000000 | 16925000000 | 16305000000 | 18925000000 | 19100000000 | 18800000000 | 19400000000 | 18900000000 | 18400000000 | 18800000000 | 18600000000 | 18200000000 | 18.5 | 18600000000 | 18500000000 | 17200000000 | 16700000000 | 17000000000 | 16500000000 | 16400000000 | 15900000000 | 15900000000 | 15300000000 | 15400000000 | 14700000000 | 14600000000 | 14300000000 | 13700000000 | 13600000000 | 13400000000 | 13500000000 | 12700000000
light dope | 22300000000 | 23500000000 | 24400000000 | 27100000000 | 29100000000 | 27800000000 | 27150000000 | 27530000000 | 27820000000 | 28200000000 | 28040000000 | 29060000000 | 28630000000 | 28870000000 | 28520000000 | 28870000000 | 28910000000 | 29000000000 | 28487000000 | 28690000000 | 28450000000 | 28460000000 | 28270000000 | 28958000000 | 28499000000 | 28460000000 | 28390000000 | 28150000000 | 28050000000
Heavy N2-dope
Baseline, no doping
Light N2-dope
2K meas.
RF surface | Qo at 16MV/m | Pi-mode limit
Heavy N2 dope (20/30) | 3.0e10 | Quench, 17MV/m
Baseline, no doping | 1.7e10 | Quench, 30MV/m
Light N2 dope (6/6) | 2.9e10 | Quench, 20MV/m

## Slide 8: Cornell, FNAL, Jlab 9 Cell Results
Average quench field:
~22 MV/m!
Conclusion: Reduced lower critical field Hc1 in N-doped cavities => earlier vortex penetration at defects => lower average quench field.
9-Cell Performance after Short (2 min) Doping

## Slide 9: Field Emission
Chart: Eacc [MV/m]
| 1.21 | 1.76 | 3.21 | 4.08 | 4.95 | 6.5 | 7.52 | 8.61 | 9.4600000000000009 | 10.64 | 11.78 | 12.79 | 14.35 | 15.73 | 18.37 | 20.29 | 19.350000000000001 | 22.89 | 25.25 | 26.84
Initial N-Dope | 22677000000 | 22622000000 | 24330000000 | 25218000000 | 26224000000 | 28908000000 | 29634000000 | 30259000000 | 29429000000 | 28831000000 | 28305000000 | 25880000000 | 23538000000 | 23432000000 | 23811000000 | 20753000000 | 20213000000 | 15525000000 | 16200000000 | 14429000000
Re-VEP | 23580000000 | 20984000000 | 25541000000 | 27425000000 | 28867000000 | 28553000000 | 27297000000 | 29805000000 | 29353000000 | 29730000000 | 30756000000 | 30112000000 | 31487000000 | 30478000000 | 31644000000 | 29850000000 | 31863000000 | 29977000000 | 31893000000 | 28772000000 | 28486000000 | 28781000000 | 28648000000 | 28735000000 | 27715000000 | 27266000000
Chart: Eacc [MV/m]
| 1.41 | 2.06 | 3.09 | 4.01 | 4.8099999999999996 | 5.93 | 7.31 | 8.19 | 9.66 | 10.66 | 11.5 | 12.29 | 13.48 | 14.78 | 17.18
Initial N-Dope | 23717000000 | 24567000000 | 23862000000 | 28542000000 | 26683000000 | 29145000000 | 32268000000 | 29764000000 | 28039000000 | 26877000000 | 25672000000 | 24006000000 | 21193000000 | 20244000000 | 17012000000
Re-HPR | 24759000000 | 22696000000 | 24861000000 | 27371000000 | 29701000000 | 31743000000 | 29763000000 | 33094000000 | 31841000000 | 32687000000 | 34027000000 | 32329000000 | 31182000000 | 33878000000 | 32576000000 | 32425000000 | 32260000000 | 31606000000 | 33129000000 | 31744000000 | 36760000000
AES018
AES022
Two of Cornell’s 9-cell cavities were limited by FE after doping
One was fixed with just an additional HPR, the other with additional VEP (which did not change Q0 performance)
Conclusion: Re-HPR and additional chemistry can be used if cavities are limited by field emission

## Slide 10: Lessons for LCLS-II
Lesson 1: Nitrogen-Doping is sufficient to reach High Q – performance can be achieved reliably at the three labs
Lesson 2: We can expect an increase of 1 to 3 nΩ of residual resistance from VT to HT
Lesson 3: Large ΔTtrans and small ΔTlong are necessary to reach High Q in N-doped cavities
Lesson 4: Pressure in the furnace does significantly not affect nitrogen uptake in cavities
Lesson 5: Great care needs to be taken with auxiliary parts such as HOM couplers, input couplers, etc.

## Slide 11: Vertical Test Procedures
Q0 vs Eacc is measured at multiple temperatures:
  Cornell measures each cavity from 1.6 to 2.1 K in 0.1 K increments
  FNAL measures each cavity at 1.5 and 2.0 K
  This allows one to decompose surface resistance in to residual and BCS portions
Cornell typically measured each cavity in a variety of cool downs with different cool down rates and external magnetic fields

## Slide 12: Vertical Test Procedures
Cavities are assembled with high Q input couplers – this gives an easy and accurate measurement of the Q0
  Cornell uses variable couplers, FNAL and JLab use fixed couplers
Vertical test dewars have ~1 mG ambient magnetic field at all three labs
Cavities are typically cooled by dumping in liquid, resulting in very fast cool downs with large temperature gradients

## Slide 13: Instrumentation and VT Summary
Item | Details
Temperature Sensors | At least 3 Cernox sensors on cavity cells
Fluxgate Magnetometers | At least two, one transverse, one longitudinal
Ambient magnetic field | <1 mG
Radiation Monitors | 1 outside the dewar under the shielding block (at Cornell)
Q0 vs E measured at | 1.6, 1.7, 1.8, 1.9, 2.0, 2.1 K (Cornell)
Other items measured | Q0 vs T, f vs T (Cornell)
Additional details | Multiple cool downs with different cooling rates and external magnetic fields (Cornell)

## Slide 14: In Production
Item | Details
Temperature Sensors | At least 3 Cernox sensors on cavity cells Need at least 1
Fluxgate Magnetometers | At least two, one transverse, one longitudinal
Ambient magnetic field | <1 mG 5 mG?
Radiation Monitors | 1 outside the dewar under the shielding block (at Cornell) Do we need more?
Q0 vs E measured at | 1.6, 1.7, 1.8, 1.9, 2.0, 2.1 K (Cornell)
Other items measured | Q0 vs T, f vs T (Cornell)
Additional details | Multiple cool downs with different cooling rates and external magnetic fields (Cornell)

## Slide 15: Cornell HTC Results
One-cavity test cryomodule
Short version of a main linac cryomodule
Same module construction (magnetic shields, cryogenic system…)
Dedicated to high Q0 studies

## Slide 16: Cornell HTC Results
Test | Cavity | Prepared By | RF Coupler | Helium Vessel | Other
HTC9-1 | ACC012 | FNAL | High Q | ILC | 
HTC9-2 | AES011 | FNAL | High Q | ILC | 
HTC9-3 | AES018 | Cornell | High Q | LCLS-II | 
HTC9-4 | AES018 | Cornell | LCLS-II | LCLS-II | 
HTC9-5 | AES030 | JLab | LCLS-II | LCLS-II | Tuner, HOM Couplers

## Slide 17: Cornell HTC Results
Cavity | Lhe Tank | HTC Test | VT Result | HT Result | ΔRVT->HT [nΩ]
TB9ACC012 | ILC | HTC9-1 | (3.5±0.4)x1010 | (3.2±0.3)x1010 | 1 ± 2
TB9AES011 | ILC | HTC9-2 | (3.4±0.3)x1010 | (2.7±0.3)x1010 | 2 ± 2
TB9AES018 | LCLS-II | HTC9-3 | (2.2±0.3)x1010 | (2.2±0.2)x1010 | 0 ± 2
Conclusion: No significant change in performance when cavity is installed in cryomodule.
2K
2K
2K
HTC9-2
HTC9-3
HTC9-1

## Slide 18: Cornell, FNAL, Jlab 9 Cell Results
Cavity ID# | 2K Q0 at 16 MV/m - vertical test, bare cavity [1E10] | 2K Q0 at 16 MV/m - vertical test, dressed cavity [1E10] | 2K Q0 at 16 MV/m - horizontal test [1E10] | Maximum accelerating field, latest test [MV/m] | DR - vertical bare to vertical dressed [nW] | DR - vertical dressed to horizontal dressed [nW]
ACC015 | 3.5 |  |  | 24.0 |  | 
AES016 | 3.0 |  |  | 20.2 |  | 
AES019 | 3.2 | 3.1 |  | 18.8 | 0.3 | 
AES021 | 3.4 | 2.8 | 3.1 | 23.0 | 1.7 | -0.9
AES022 | 3.1 |  |  | 26.2 |  | 
AES024 | 3.2 | 3.2 |  | 22.0 | 0.0 | 
AES026 | 2.8 | 2.8 |  | 21.4 | 0.0 | 
AES027 | 3.6 | 2.7 | 2.8 | 22.8 | 2.5 | -0.4
AES028 | 3.5 | 3.0 |  | 23.0 | 1.3 | 
AES029 | 3.6 | 3.6 |  | 23.7 | 0.0 | 
AES030 | 2.9 | 2.5 |  | 18.2 | 1.5 | 
AES031 | 3.5 |  | 2.4 at 8 MV/m | 19.4 |  | 
AES032 | 4.2 | 2.8 |  | 23.0 (admin limit) | 3.2 | 
AES033 | 3.9 | 3.6 |  | 21.3 | 0.6 | 
AES034 | 3.9 | 3.5 |  | 22.5 | 0.8 | 
AES035 | 3.6 | 2.9 | 3.0 | 17.5 | 1.8 | -0.3
AES036 | 4.1 | 3.7 |  | 19.0 (admin limit) | 0.7 | 
Average | 3.5 | 3.1 | 3.0 | 21.6 | 1.1 | -0.5
Conclusion: We see a 1 to 3 nΩ increase in residual resistance from bare VT to dressed HT
Should we increase the spec to 3.4x1010 in VT to meet 2.7x1010 in CM?
Notes: Add prediction to VT spec to reach 2.7 HT

## Slide 19: Decomposition of Rs
In HTC9-3, both BCS resistance and residual resistance changed from VT to HT
Should we be taking 1.6 and 2 K curves during production to identify which component of Rs potential issues are coming from?

## Slide 20: HT Procedures
Q0 vs Eacc is measured cryogenically due to low Qext, which is very time consuming
Cool downs are typically done by dumping in helium gas, which results in fast cool downs but with smaller temperature gradients than can be achieved in vertical test
Multiple cool downs completed to optimize flux expulsion and reach high Q0

## Slide 21: Instrumentation and HT Summary
Item | Details
Temperature Sensors | Cernox sensors distributed on the cavity cells, beam tubes, coupler, and HOM cans
Fluxgate Magnetometers | At least two, one transverse, one longitudinal
Ambient magnetic field | <3 mG
Radiation Monitors | On either side of the cryomodule
Q0 vs E measured at | 1.6, 2.0 K
Other items measured | Q0 vs T, f vs T (Cornell)
Additional details | Multiple cool downs with different cooling rates and external magnetic fields (Cornell)

## Slide 22: In Production
Many Cernox sensors are impractical from a cost perspective – how many should we have?
Fluxgates are also expensive but we need at least one to measure magnetic fields
1.6 K and 2.0 K curves should be measured in order to better understand poor performance
  Some cavities will not meet spec, understanding why is very important to solving that problem in future cavities
Temperature sensors on the HOM cans are highly recommended – will be discussed in detail later

## Slide 23: Lessons for LCLS-II
Lesson 1: Nitrogen-Doping is sufficient to reach High Q – performance can be achieved reliably at the three labs
Lesson 2: We can expect an increase of 1 to 3 nΩ of residual resistance from VT to HT
Lesson 3: Large ΔTtrans and small ΔTlong are necessary to reach High Q in N-doped cavities
Lesson 4: Pressure in the furnace does not significantly affect nitrogen uptake in cavities
Lesson 5: Great care needs to be taken with auxiliary parts such as HOM couplers, input couplers, etc.

## Slide 24: Magnetic Field Studies
Temperature Sensors
Fluxgate
Magnetometer
Helmholtz
Coil
Slow Cool
Down System

## Slide 25: Magnetic Field Studies
Stronger doping results in a higher sensitivity to trapped flux.
Nitrogen-doped cavities showed a higher sensitivity than EP and EP+120oC baked cavities.

## Slide 26: Cool Down Studies
Rres vs ΔTvert
Rres vs Helium Mass Flow Rate
Conclusion: Helium flow rates of >2 g/s needed for efficient magnetic field expulsion by vertical temperature gradients.

## Slide 27: Cool Down Studies
i1
i2
Cavity primarily cools from bottom to top => large ΔTvertical in fast cool down
  Good for efficient magnetic field expulsion
  But: conductivity  = (T) => Cylindrical symmetry is broken!
    Finite ΔThorizontal will drive thermal-electric currents with preferential flow through the bottom of the cavity
    Non-zero magnetic field at the cavity inner surface
Ideal cool-down: large ΔTvertical by fast cool down with small ΔThorizontal
DTvertical
DThorizontal
D. Gonnella et al., J. Appl. Phys. 117 , 023908 (2015)
R. Eichhorn, arXiv:1411.5285 [physics.acc-ph] (2014)
J.-M. Vogt et al., Phys. Rev. ST Accel. Beams 18, 042001 (2015)
Nb
Ti
Nb
Ti

## Slide 28: Cool Down Studies
Btrans vs ΔThoriz
Rres vs ΔThoriz
Conclusion: Small horizontal temperature gradients <10K critical to keep impact of thermal-electric currents on Rres small.
The LCLS-II Helium vessel helps to minimize ΔThoriz

## Slide 29: Lessons for LCLS-II
Lesson 1: Nitrogen-Doping is sufficient to reach High Q – performance can be achieved reliably at the three labs
Lesson 2: We can expect an increase of 1 to 3 nΩ of residual resistance from VT to HT
Lesson 3: Large ΔTtrans and small ΔTlong are necessary to reach High Q in N-doped cavities
Lesson 4: Pressure in the furnace does not significantly affect nitrogen uptake in cavities
Lesson 5: Great care needs to be taken with auxiliary parts such as HOM couplers, input couplers, etc.

## Slide 30: Diffusion Simulation
Good agreement between measures interstitial nitrogen doping profile and diffusion model prediction.
We have developed a diffusion simulation code that predicts nitrogen concentration in niobium based on doping parameters

## Slide 31: Diffusion Simulation
Measured furnace pressure vs. time
Drop in furnace pressure vs. time
N2 gas injections into furnace
Nitrogen uptake during doping is not dependent on pressure
Conclusion: Exact pressure in the furnace does not matter to achieve the same doping level

## Slide 32: Lessons for LCLS-II
Lesson 1: Nitrogen-Doping is sufficient to reach High Q – performance can be achieved reliably at the three labs
Lesson 2: We can expect an increase of 1 to 3 nΩ of residual resistance from VT to HT
Lesson 3: Large ΔTtrans and small ΔTlong are necessary to reach High Q in N-doped cavities
Lesson 4: Pressure in the furnace does not significantly affect nitrogen uptake in cavities
Lesson 5: Great care needs to be taken with auxiliary parts such as HOM couplers, input couplers, etc.

## Slide 33: Impact of HPC on Q0
2K Q0,cryo vs. RF power (under full reflection)
2K Q0,cryo vs. Qext
Conclusion: No significant increase in 2K cryogenic load or cavity performance degradation from RF input coupler.
HTC9-4: RF power coupler
HTC9-3: high Q antenna

## Slide 34: HOM Can Multipacting
HOM can temp.
Short (due to HOM can fabrication error) and multipacting in one HOM coupler resulted in significant heating and Q-slope
FNAL also observed this behavior but was able to condition it away

## Slide 35: Lessons for LCLS-II
Lesson 1: Nitrogen-Doping is sufficient to reach High Q – performance can be achieved reliably at the three labs
Lesson 2: We can expect an increase of 1 to 3 nΩ of residual resistance from VT to HT
Lesson 3: Large ΔTtrans and small ΔTlong are necessary to reach High Q in N-doped cavities
Lesson 4: Pressure in the furnace does not significantly affect nitrogen uptake in cavities
Lesson 5: Great care needs to be taken with auxiliary parts such as HOM couplers, input couplers, etc.

## Slide 36: Lessons for LCLS-II
Lesson 1: Nitrogen-Doping is sufficient to reach High Q – performance can be achieved reliably at the three labs
Lesson 2: We can expect an increase of 1 to 3 nΩ of residual resistance from VT to HT
Lesson 3: Large ΔTtrans and small ΔTlong are necessary to reach High Q in N-doped cavities
Lesson 4: Pressure in the furnace does not significantly affect nitrogen uptake in cavities
Lesson 5: Great care needs to be taken with auxiliary parts such as HOM couplers, input couplers, etc.

## Slide 37: Main Takeaway
In every vertical and horizontal test, we encountered something unexpected – it is reasonable to assume that we will continue to experience this at the beginning of production